## Slide 1
1
Calibration Adjustments for the MODIS Aqua 2015 Ocean Color Reprocessing
Gerhard Meister, NASA Code 616
OBPG (Ocean Biology Processing Group)
5/18/2015
MODIS/VIIRS Science Team, Calibration Breakout
Silver Springs, MD,
May 2015
1
Notes: 1
1
1
1
monitoring harmful blooms from space gives coastal communities ability to take precautionary measures
  fishing bans
  beach closures
image shows true color (top) and chlorophyll (bottom) views of bloom in mouth of Mississippi
similar information was used to monitor a red tide bloom off coast of Texas last fall

## Slide 2
2
Background:
The upcoming MODISA Ocean Color reprocessing (R2014.0) is mainly about a new data format (HDF5) for the L2 files, ancillary data, algorithm updates
We used new fits for the whole time series (avoids uncertainty from predictions during regular processing)
MCST provided complete refit of RSB LUT (now V6.1.33.8_OC2)
OBPG redid xcal analysis (now axc38d) of MODISA to itself
MODIS Aqua does not show strong changes in polarization sensitivity -> polarization correction still prelaunch, xcal adjustments only for gains
2
Notes: 2
2
2
2

## Slide 3
Xcal: MODISA ‘true’ TOA radiance
Lt(λ) = [ Lr(λ) + La(λ) + tLf(λ) + TLg(λ) + td(λ)Lw(λ) ] · tg(λ)
air
aerosol
whitecap
glint
water
gas
from MODISA NIR
assumes MCST NIR band characterization
from MODISA Lw(λ) 6-day mean
(7 days excluding central day,
only frames 300-1050)
L(λ) = radiance at wavelength λ
t (λ) = transmission at wavelength λ
Xcal to itself: small impact on temporal trending, potentially large impact on scan angle depencence

## Slide 4
Crosscalibration coefficients:
Lm/M11 = Lt
M11= a0+a1*f+a2*f2 +a3*f3 +a4*f4
Lm: measured TOA radiance (MODISA)
Lt: true TOA radiance (from MODISA L3)
f: frame number (0-1353)
M11: gain correction, depends on band, MS, detector. Total of 2*10*5 = 100 polynomial coefficients (a(i)) per band per day. ~2-3 million matchups per day.
4
Notes: 4
4
4

## Slide 5
M11 at 412nm for 4 different scan angles:
5
Colors: detectors 1-10
Solid (dashed) lines: mirror side 1 (2)
Beg. of scan
(lunar)
Nadir
Solar Diffuser
End of scan
Notes: 5
5
5

## Slide 6
M11 at 547nm for 4 different scan angles:
6
Colors: detectors 1-10
Solid (dashed) lines: mirror side 1 (2)
Beg. of scan
(lunar)
Nadir
Solar Diffuser
End of scan
Notes: 6
6
6

## Slide 7
M11 at 547nm after temporal fit:
7
Colors: detectors 1-10
Solid (dashed) lines: mirror side 1 (2)
Beg. of scan
(lunar)
Nadir
Solar Diffuser
End of scan
Notes: 7
7
7

## Slide 8
M11 scan angle dependence:
Average (black), BOM (blue), EOM (red)

## Slide 9
Temporal anomaly: 412nm Rrs
Operational:
Reprocessing:
~5% downward trend in reproc. over 13 years
Operational end of mission artifacts due to not updating neither MCST LUT nor xcal

## Slide 10
Temporal anomaly: 443nm Rrs
Operational:
Reprocessing:
No apparent trend in reproc. over 13 years, higher variability at end and beginning
Operational end of mission artifacts due to not updating neither MCST LUT nor xcal

## Slide 11
Temporal anomaly: 547nm Rrs
Operational:
Reprocessing:
Slight upward trend in reproc., stronger in olig., less in meso. and eutr.
Operational end of mission artifacts due to not updating neither MCST LUT nor xcal

## Slide 12
Temporal anomaly: 667nm Rrs
Operational:
Reprocessing:
No apparent trend in reproc., same for 678nm
Operational had a clear downward trend (about 10%), same for 678nm

## Slide 13
Temporal anomaly: Chl-a, oligotrpohic
Operational:
Reprocessing:
Chlorophyll-a in oligotrophic water is ~10% higher in 2014 for reprocessing

## Slide 14
Temporal anomaly: Chl-a, mesotrpohic
Operational:
Reprocessing:
Chlorophyll-a in mesotrophic water is ~15% higher in 2014 for reprocessing

## Slide 15
Scan angle dependence of 412nm nLw (reproc.):
Early mission:
Late mission:
Increased variation around frame 900 (glint region)
Increased striping and overall variability towards end of mission

## Slide 16
Scan angle dependence of 547nm nLw (reproc.):
Early mission:
Late mission:
Increased variation around frame 900 (glint region)
No (global average) striping, not much degradation with time

## Slide 17
Scan angle dependence of chl-a (reproc.):
Early mission:
Late mission:
Increased variation around frame 900 (glint region)
Striping and scan angle dependence increase towards end of mission

## Slide 18
18
Summary:
Reprocessing improves the most recent time period, but variability in blue bands remains
Overall improvement for long term trend of red bands, but 412nm long term trend still questionable
Chlorophyll-a ~10-15% higher in 2014 after reprocessing
Using MODIS Aqua results, MODIS Terra will be crosscalibrated in the next few months (SeaWiFS and MERIS first). Switch to new data format will be made soon without recalibration. Time series may be discontinuous e.g. due to new ancillary data until new crosscalibration is applied.
18
Notes: 18
18
18
18

## Slide 19
Backup slides:

## Slide 20
Xcal approach (Step 1): input ‘true’ L3 (15 days) Rrs
20
MODIS Aqua: 412nm Rrs (remote sensing reflectance)
Notes: 20
20
20

## Slide 21
Xcal approach (Step 2): input MODIS Lm
21
MODIS Terra TOA radiance from one day, screened for chlorophyll, AOT, straylight, etc.
Notes: 21
21
21

## Slide 22
Xcal approach (Step 3): vicarious Lt (from ‘true’ L3)
22
Lt(λ) = [ Lr(λ) + La(λ) + tLf(λ) + TLg(λ) + td(λ)Lw(λ) ] · tg(λ)
Notes: 22
22
22

## Slide 23
Xcal approach (Step 4): Ratio Lm/vLt
23
Lm/M11 = Lt + m12*Q + m13*U
Notes: 23
23
23