## Slide 1
1
S-NPP, N-20, and N-21 VIIRS Reflective Solar Bands
On-orbit Radiometric Calibration and Performance
VIIRS Characterization Support Team (VCST), NASA GSFC
(presented by Ning Lei)
May 1, 2023

## Slide 2
2
Objectives
Radiometric calibration improvements since last STM (Feb. 2021)
RSB radiometric performance update
Future works

## Slide 3
3
What happened since STM 2021
Safe mode occurred in July 2022 led to F-factor trending change for many
RSBs
Developed TOA reflectance factor uncertainty algorithm
Improved L1B pixel saturation detection
S-NPP
Sliding window (in time) approach, using F(moon): same approach employed by S-NPP
Improved L1B pixel saturation detection
N-20
Launched on Nov. 10, 2022
First mission RSB F-factor LUTs delivered with screen functions derived
from calibration data collected on yaw maneuver orbits
N-21

## Slide 4
4
VIIRS RSB calibration
Space view
SD screen
SDSM screen
sunlight
SDSM
RTA
Solar Diffuser (SD)
Solar Diffuser (SD): a calibration source; its BRDF change (H-factor) measured by the SD stability monitor (SDSM)
on-orbit calibrated
Spectral radiance

## Slide 5
5
Performance of S-NPP SD and SDSM
SDSM detector gains trend normally
H-factors dropped by about 1% due to the Feb. 24, 2019 event
Recently, H-factors trend upwards (det 1-5), because of unusual solar activity impact on SD, also occurring for N20 and MODIS SDs

## Slide 6
6
Calculate F(SD)
Solar azimuth angle
and
obtained from fitting F(SD) to F(Moon)
S-NPP and N-20
N-21
Not enough F(Moon) available to fit; choose to use the measured HSDSM (no deconvolution)
Use a sliding window (in time) approach, fit F(HRTA) to F(Moon) to find

## Slide 7
S-NPP RSB 1/F-factors
Solid lines: 1/F(SD); Circles: 1/F(Moon)
F decreases with time, because of telescope mirror surface tungsten oxide contamination

## Slide 8
N-20 H-factors
H-factors decrease at smaller rates than S-NPP H-factors
SDSM gains decrease similarly to S-NPP
Performance of N-20 SD and SDSM
N-20 SDSM Detector Gains
Notes: 8
SDSM_Hplot_V4aeroX.20150629

## Slide 9
N-20 RSB 1/F-factors
F-factors are very stable over mission
Solid lines: 1/F(SD); Circles: 1/F(Moon)

## Slide 10
N-21 H-factors
H-factors decrease faster than N-20 and similarly to S-NPP H-factors
SDSM gains decrease similarly to S-NPP and N-21
Undulation in the H mainly from screen function errors (will be improved
later this year)
Performance of N-21 SD and SDSM
N-21 SDSM Detector Gains
Notes: 10
SDSM_Hplot_V4aeroX.20150629

## Slide 11
Performance of N-21 RSBs: F-factors
VISNIR bands: F(SD, with measured HSDSM ) has a downward trend
SWIR bands: F(H=1) trends upwards, likely because of ice accumulation on focal plane
Mid-mission outgassing

## Slide 12
SNRs
All SNRs satisfy requirements, except N-20 I3 band detector 29 (noisy detector)

## Slide 13
13
S-NPP H-factor SD positional dependence
SD BRDF change factor (H-factor) can depend on SD position
detector array
model parameters

## Slide 14
M1 striping gone (C2.0)
M1 striping (C1.1)
2019229
H-factor SD positional dependence algorithm (2017) and results (2018) are still good
S-NPP TOA images
March 2023
2019229
No striping seen (C2.0)

## Slide 15
Uncertainty of TOA reflectance factor
Derive reflectance factor uncertainty from definition
Reference: VCST_TECH_REPORT_2022_009
+
+ +
Uncertainty LUTs are generated for L1B code for S-NPP and N-20

## Slide 16
Improvement in detecting saturation/rollover
Original rollover flagging (black dots)
Current rollover flagging (black dots)
Use neighboring band radiance as a reference, in additional to its own radiance
Reference: A. Angal et al., SPIE Proc. Vol. 12232 (2022)

## Slide 17: Future works
  1. N-20 RSB striping mitigation
  2. Improve N-21 screen functions with data collected at both yaw
  maneuvers and regular orbits
  3. S-NPP M6 saturation flagging

## Slide 18: Summary
S-NPP, N-20, and N-21 VIIRS RSBs perform normally with SNRs satisfying specifications (except N-20 I3 detector 29) and will remain above specifications for the foreseeable future
Used N-20 RSB F(moon) to find HRTA from HSDSM, instead of using S-NPP HRTA results, a sliding window approach
Developed L1B reflectance uncertainty algorithm, delivered uncertainty LUTs
Improved saturation detection algorithm